INTERCOMPARISON OF LABORATORY MEASUREMENTS OF AIRBORNE SOUND INSULATION OF PARTITIONS

A. Farina (1), P. Fausti (2), R. Pompoli (2), F. Scamoni (3)

(1)	Dipartimento di Ingegneria Industriale, University of  Parma

(2)	Istituto di Ingegneria, University of  Ferrara

(3)	ICITE - Istituto Centrale Industrializzazione e Tecnologia Edilizia (MI)

(INCE=33)

# INTRODUCTION

Since it is not possible to specify completely the construction of laboratory test facilities or the sound field conditions some details of the test facilities and procedures are left to the choice of the operator. This, together with the statistical character of sound fields within rooms, leads to uncertainties in the results due to non-systematic and systematic influences. In agreement with modern statistical methods, the concepts of repeatability  “r” and reproducibility “R” have been used in ISO 140-2 [1] to state the precision of the measurements carried out according to a test method. Repeatability “r” shows the closeness of agreement between mutually independent test results obtained with the same method on identical test material in the same laboratory with the same equipment by the same operator within short intervals of time. Reproducibility “R” shows the closeness of agreement between test results obtained with the same method on identical test material in different laboratories with different operators using different equipment. The determination of the repeatability and reproducibility values of a test method is obtained by an interlaboratory test taking into account the procedures given in the international standards ISO 140-2 [1] and  ISO 5725 [2]. Tentative values of “r” and “R” are given in Annex A of ISO 140-2. The “r” and “R” values may also be used to verify the proper operation of test procedures of a laboratory which has not taken part in the intercomparison.

With the aim to determine the values of “r”  and “R”  for the sound reduction index of walls, the Engineering Institute of the University of Ferrara, with the support of SINAL (National System for Laboratory Accreditation) and the GAA (Environmental Acoustic Group) of the Italian Acoustic Association organized an interlaboratory test among nine Italian laboratories.

# PARTICIPANT LABORATORIES

Nine Italian laboratories took part in this intercomparison: “C.S.I” in Bollate (MI), “C.S.R. - Consorzio Studi e Ricerche della SCM” in Rimini, “I.C.I.T.E. - Istituto Centrale Industrializzazione e Tecnologia Edilizia” in S.Giuliano Milanese (MI), “Istituto di Acustica O.M.CORBINOÓ in Rome, “I.E.N.G.F. - Istituto Elettrotecnico Nazionale G. Ferraris” in Turin, “ISTITUTO GIORDANO - Centro Studi e Ricerche di Fisica Tecnica” in Bellaria (RN), “ISTEDIL - Istituto Sperimentale per l'Edilizia” in Guidonia Montecelio (ROME), “MODULO UNO - Acoustical Engineers and Consultants” in Turin, “SIMPA” in Ferentino (FR). Seven of these laboratories certify the sound insulation performance of walls for external customers, two of them use the laboratory for the qualification of their own products or for research activities. It seems important to note that the laboratories which took part in this intercomparison were not chosen after verifying the conformity of their test facilities with ISO 140-1, but simply because they were available and  interested in participating in these tests. Table 1 shows the main features of these laboratories.

Table 1: Main features of the nine laboratories

| Laboratories | Corbino | Csi | Csr | Icite | Iengf | Istedil | Istituto Giordano | Modulo Uno | Simpa |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Dimensions room 1 (m) | 4.8x 5.84x 3.05 | 3.7x 4.05x 3.35 | 4.45x 5.7x 3.06 | 4.4x 4.4x 3.20 | 3.97x 4.3x 3.46 | 4.12x 5.0x 3.18 | 3.62x 4.4x 3.50 | variable | 3.8x 2.9x 4.71 |
| Volume room 1 (m3) | 85.5 | 50.2 | 74 | 53.6 | 59 | 65.5 | 56 | 76 | 52 |
| Dimensions room 2 (m) | 4.8x 5.84x 3.05 | 6.0x 4.05x 3.35 | 4.15x 5.0x 3.25 | 4.85x 4.0x 3.70 | 4.1x 4.3x 3.46 | 4.12x 5.0x 3.45 | 4.7x 4.4x 3.50 | variable | 3.8x 2.9x 4.95 |
| Volume room 2 (m3) | 85.5 | 81.4 | 67.5 | 59.3 | 65 | 71.1 | 72 | 222 | 54.5 |
| Test opening (m) | 2.5x 2.54 | 4.05x 2.7 | 2.0x 2.2 | 3.31x 3.0 | 3.6x 2.8 | 3.1x 2.96 | 3.6x 3.0 |  | 3.6x 2.74 |
| Area test opening (m2) | 6.3 | 10.9 | 4.4 | 10 | 10.08 | 9.23 | 10.8 | 9.2 | 9.86 |
| Diffusing elements | no | no | si | no | si | si | si | si | no |
| Rw max (ISO 717) |  |  | 56.5 |  | 61 | 60 |  | 58 | 56 |

# THE SPECIMENS

Two different gypsum board partitions were tested in each laboratory. The first wall had one single plate of gypsum board, with a thickness of 12,5 mm, on each side. The plates were mounted on a metal frame made by studs and guides 75 mm thick. Glass wool 50 mm thick and with a density of 40 kg/m3 was inserted inside the wall. The second specimen was a wall with two gypsum boards on each side. The frame was as described for the first specimen. The two specimens were supplied free by BPB Italia, which also provided a single team of workers for mounting the walls in all the laboratories. The material came from one single stock of production and the mounting was performed according to a well defined procedure.

# MEASUREMENTS AND RESULTS

All the measurements were carried out according to the international standard ISO 140-3 (1995) [3]. This standard introduces many changes with respect to the previous version of ISO 140-3 (1978) and its amendment (1990). It gives more details for the installation of partitions, doors, windows, glazing and facade elements. The generation of the sound field in the source room has to comply with stricter conditions. Special procedures have to be followed for the qualification of the loudspeakers and of the loudspeaker positions. The frequency range of measurements has been extended to the centre frequency of 5000 Hz and additional information in the low frequency range ( 50, 63, 80 Hz) is also required.

The tests started at the end of 1993 and finished at the beginning of 1995. Each laboratory was indicated with a letter (A-I) in order to ensure anonymity. Not all the laboratories could supply the 5 independent tests required by ISO140 -2: in this case the values of the simple variance ÒsrÓ were assigned on the basis of previous measurements taken in the same laboratory with other specimens.

From the analysis of all the data, laboratory ÒIÓ came out as anomalous and therefore its measurements were not used for computing the “r” and “R” values.

Figure 1 shows, for the two specimens, the sound reduction index as a function of the frequency in the range 100 Hz - 3150 Hz supplied by all the laboratories. The frequency range was limited to 3150 Hz as laboratory ÒHÓ did not give results at frequencies higher than this value.

Figure 2 shows, for the two specimens, the values of the single number of the sound reduction index Rw (dB) computed as required by ISO 717  [4]. The first specimen gave Rw=46 dB; the second specimen gave Rw=51 dB. The same figure 2 shows two more single numbers of the sound reduction index  indicated as dB(A) pink and dB(A) traffic obtained according to the last revision of ISO 717.

Figure 3 shows, for each specimen, the values of “r” and “R” obtained from this interlaboratory test. These values are compared with the tentative values given in Annex A of ISO 140-2.

# CONSIDERATIONS

The repeatability values obtained in this interlaboratory test (see figure 3) are lower than the values of Annex A. This result is probably due to the automatic measurement procedures made possible by modern instrumentation, which reduce the causes of variabilitiy for results obtained in a laboratory which operates under repeatability conditions. On the contrary, the reproducibility values are higher than the values of Annex A (fig.3).

Figure 1: Sound reduction index of specimen 1 and 2 obtained from all the measurements.

Figure 2: Single-numbers quantities of sound reduction index of specimens 1 and 2: Rw, Rpink, Rtraffic.

It seems important to note that these “R” values, quoted in Annex A of ISO 140-2, are tentative and represent an average  of “R” values obtained from three different interlaboratory tests. One was carried out among 5 laboratories in Germany in 1983, using as test specimen a glazing in a staggered test opening; the second was an interlaboratory test conducted from 1982 to 1985 involving seven laboratories in Belgium and in the Netherlands using a lightweight partition and two brick walls; the third was an interlaboratory test conducted in 1985/86, a BCR project involving eight laboratories using double-glazing. ISO 140-2 does not give the dispersion of the data used for the R values shown in Annex A. The lack of sufficient data and the difference in  these experiences do not make it possible to draw a significant comparison with the results obtained under this interlaboratory test.

Figure 3: Repeatability r and reproducibility R values obtained from this interlaboratory test: comparison with the values given in ISO 140-2.

A critical analysis of all the results has not yet been done. However, some observations are possible and some hypotheses may be formulated to interpret the data.

The distribution of “R” with the frequency is irregular, and shows evident maxima and minima (fig.3). Specimen 1, the less insulating, shows a maximum value at 400 Hz  (R =12 dB) and a minimum at 1600 Hz (R =4 dB); in the low frequency range, the R  values are lower than the ones in Annex A. The reproducibility value of Rw is 4 dB. Specimen 2, more insulating, shows two maxima at 200 Hz and at 400 Hz with R =12 dB, and a minimum at 2000 Hz with R = 6 dB.  Generally, the “R” values are higher than those of specimen 1, as is also evident from the reproducibility value of Rw which is 4,5 dB. The increase in the reproducibility values with the sound insulation performance of the specimens shows a possible effect of flanking transmission between the test rooms in the laboratories which took part in this intercomparison. There are also other factors which may have produced high reproducibility values: the differences among the specimens mounted in the laboratories due to the different test opening dimensions (see table 1) and to the unavoidable differences in the mounting, even if this was done by one single team of workers operating under precise guidelines. The conditions of sound diffusion in the test rooms are different between the laboratories, as is evident from the significant variations of the reverbaration time values. In some laboratories the reverbaration times were much higher than the values suggested by the international standard ISO 140-1 which gives the requirements for the laboratory test facilities.

# CONCLUSIONS AND FUTURE PROJECTS

These results show that measurements can be influenced by many factors, like mounting and the characteristics of the test specimen, the dimension and shape of the laboratories and test openings, flanking transmission, diffusion of sound field, reverberation time, measurements procedures and instrumentation. The limited number of laboratories did not make it possible to carry out a comparison between laboratories with similar acoustic and geometric characteristics. Howewer, the comparison was useful in order to determine the parameters which need to be strictly controlled and to provide information for the future revision of the national and international standards.

On the basis of the experience acquired with this interlaboratory test, a comparison between European laboratories, financially supprted by the European Community (Measurements and Testing Programme 1990-94) was organized. The intercomparison started in 1995 and will be finished in the spring of 1997. Twenty-four laboratory from eleven different countries are taking part in the project.

Acknowledgments

The Authors wish to thank all the participants for their collaboration. Special thanks to BPB Italia for supplying and mounting the test specimens.

References

[1] ISO 140/2 (1991), ÒMeasurement of sound insulation in buildings and of building elements - Part 2: Determination, verification and application of precision dataÓ, UNI EN 20140-2 (1994).

[2] ISO 5725 (1986), ÒPrecision of test methods - Determination of repeatability and reproducibility for a standard test method by inter-laboratory testsÓ.

### [3] ISO 140/3 (1995), ÒMeasurement of sound insulation in buildings and of building elements - Part 3: Laboratory measurements of airborne sound insulation of building elementsÓ.

### [4] ISO/DIS 717/1-2 (1994), ÒRating of sound insulation in buildings and of buildings elements. Part 1: Airborne sound insulation in buildings and of interior building elements. Part 2: Impact sound insulation”.

[5] ISO/DIS 140/1 (1994), ÒMeasurement of sound insulation in buildings and of building elements - Part 1: Requirements for laboratory test facilities with suppressed flanking transmissionÓ.